## Slide 1: EPIC UV Channel Calibration Update
DSCOVR EPIC and NISTAR Science Meeting
Sept 27-29, 2022
Liang-Kang Huang

## Slide 2: Lunar Measurements for Sensitivity Dependence as a function of pixel distance to CCD center
EPIC UV channel flat field calibration corrections were derived in 2019 with comparison of accumulated measurements over the 18 months between EPIC and SNPP OMPS NM.
The lunar data can be used to resolve uncertainty in the UV channel flat field calibration, which is due to RTM error at high zenith angles and is expected to be circular symmetric.
Dedicated EPIC lunar measurements were performed in May 2021 and September 2021, which covered all 10 EPIC channels. Then, the measurement sequence is repeated in July 2022.
The lunar data were analyzed. The sensitivity changes as a function of pixel distance to CCD center were estimated for every wavelength channel.

## Slide 3
DSCOVR EPIC Lunar Measurements in 2021 & 2022
Back ground rim
For Radial Dependence in Flat Field Calibration
Moon image centers
⋇ in 2021 ▢ in 2022
Large symbols were planned positions and small symbols were actual positions.
Lunar phase changes

## Slide 4
May 2021
July 2022
July 2022
Sept 2021

## Slide 5
Repeatable ~0.5%
Measurements in 2021 & 2022 are fitted together
Results differ channel by channel
Eight channels have significant lower values near the edge.
Variations with 680 nm & 688 nm channels are less than uncertainties in the measurements.
More lunar measurements are needed to improve accuracy.
The results are recommended for the flat field calibration correction in L1 data processing.
Results from Lunar Measurements

## Slide 6: EPIC UV Channel Albedo Calibrations Coincidence Comparison with SNPP OMPS NM
EPIC v03 L1a and NM AS61006 NMEV-L1B are used.
EPIC v03 L1a counts are also corrected for the sensitivity change from the CCD center to the CCD edge, based on the 2021 & 2022 lunar data.
Radiance measurements from the two instruments were taken mostly within 10-15 minutes.
Selected EPIC pixels filled in NM’s field of view at the same geolocation.
Solar zenith angle difference between EPIC and NM in the comparison is less than 2o.
Backscattering angle difference between EPIC and NM is less than 1o.
NM albedo spectrum (radiance normalized with solar flux) are interpolated to EPIC wavelengths, except EPIC 388 nm channel which is compared with NM 380 nm.
Viewing geometry differences are simulated with RTM TOMRAD, and corrected.

## Slide 7
The ratio is equal to EPIC’s radiance calibration coefficient.
Dots are raw ratio measurements.
The data are smoothed with a piece-wise linear fit with 7-day interval.

## Slide 8
There are annual cycles and ~6-month cycles. The 6-month cycles may be related to EPIC’s orbit rotation.
Raw ratio measurements are fitted with a cubic polynomial.
Slope at 2021.25 is used to for a constant drift rate to cover time after year 2021.25.
More data after May 2022 are needed to have complete cycles for the current drift rate estimate.

## Slide 10: EPIC UV Channel Radiance Calibration Coefficients
A0 values are the radiance calibration coefficients at t=2016, which are almost identical to the previous values.
A(t) = A0 * ACF(t) will give you time dependent radiance calibration coefficients.
ACF(t) are the tabulated time dependence from the fitted color curves, which are normalized at t=2016.
Flux values are at AU 1 distance. EPIC albedo calibration must use the given AU1 flux values in order to be consistent with the albedo calibrations of SNPP OMPS NM.
Channel | λ (nm) | A0 | σA0 | Flux (AU1)
1 | 317.478 | 3.199e-02 | 1.904e-06 | 819.00
2 | 325.035 | 2.887e-02 | 1.879e-06 | 807.69
3 | 339.858 | 6.357e-03 | 4.846e-07 | 995.81
4 | 387.923 | 8.722e-03 | 9.237e-07 | 1003.33

## Slide 11: backup